## Title
<<INDEX
|  |  | Project Name | ClotQuant 2
|  |  | Document Details | ClotQuant 2 Bringup Test
|  |  | Version | 2
Revision History
| Sr No | Date | Revision | Release Notes | Author
| 1 | 2025-03-05 | 1 | Draft Version | SqM Technologies
| 2 | 2025-05-09 | 2 | Draft Version | SqM Technologies
<<INDEX

## ClotQuant2-BD
=DISPIMG("ID_996526CFEDAF4D1C933510E39AFAE41B",1) | Connector Mapping

## Connection
Connector Details - 1 | Connector Details - 1

## Index
<<Title
ClotQuant2 Board Bring up (CQ2BB) Report
| SR. No. | Test Group | Test Cases | Status
| 1 | NA | Test Setup | NA
| 2 | NA | Design History | NA
| 3 | CQ2BB1 | Visual Inspection | PASS
| 4 | CQ2BB2 | Impedance Measurement | PASS
| 5 | CQ2BB3 | Supply Rails | PASS
| 6 | CQ2BB4 | SWD LED code flash | PASS
| 7 | CQ2BB5 | Temperature Sensor | PASS
| 8 | CQ2BB6 | Display Test | PASS
| 9 | CQ2BB7 | Pipette | PASS
| 10 | CQ2BB8 | WiFi module | PASS
| 11 | CQ2BB9 | BLE | PASS
| 12 | CQ2BB10 | USB OTG 1 | PASS
| 13 | CQ2BB11 | USB OTG 2 | PASS
| 14 | CQ2BB12 | RFID | PASS
| 15 | CQ2BB13 | Printer | PASS
| 16 | CQ2BB14 | Ethernet | PASS
| 17 | CQ2BB15 | Heater | PASS
| 18 | CQ2BB16 | SDRAM | PASS
| 19 | CQ2BB17 | QSPI NOR | PASS
| 21 | CQ2BB18 | ADC Read | PASS
| 22 | CQ2BB19 | IR LED test | PASS
| 23 | CQ2BB20 | RS232 | PASS
| 24 | CQ2BB21 | Buzzer | PASS
| 25 | CQ2BB22 | FAN | PASS
| 26 | CQ2BB23 | RTC | PASS
| 27 | CQ2BB24 | Encryption IC | UNATTEMPTED
| 28 | CQ2BB25 | J11 Connector | UNATTEMPTED
<<Title

## Test Setup
<<INDEX
Board and Test Details
| Board Serial Number | Version3.0 No.2
| Test Engineers | Harshit Shakya
Bare Board Received Date
Assembled Board Received Date
Date Of Testing
List of Equipment
| Description | Quantity | Remark
| Power Supply | 1
| DMM | 1
| Oscilloscope | 1
<<INDEX

## Design History
<<INDEX
Design History
List Of DNI Components
Reference Designator
L6
R27, R33, R64, R68, R81, R105
U26 (Due to internal Vref of ADC)
List Of HW Changes
| Reference Designator | Date | Change Description | Images | Remark
| U18 | 2025-04-14 | Connect U18 Pin33 (IRQN) to MCU PB1 pin. | =DISPIMG("ID_3D7D83ABAF524CD7902A430F61513E27",1) | Changes has been done.
| J7 | 2025-04-14 | Printer Power J7 pin 3 requires a 5V supply. | =DISPIMG("ID_E33C4A193FD84952A980861C4B8DD1E7",1) | Changes has been done.
| R22, R25,R47,R96 | 2025-04-14 | IR receiver op-amp feedback resistor R22, R25,R47,R96 should always be 5K or below. | =DISPIMG("ID_E64239F61E8C4EE4B55CFBEC8FD1F842",1) | Have to set manually.
| R7 | 2025-04-14 | Pipette R7 routing is missing for 3.3V. | =DISPIMG("ID_813D6950A26B407C8D59F770CAB74333",1) | Changes has been done.
|  | 2025-04-14 | Connect 12VB to the 12V pin. | =DISPIMG("ID_EEBB852DB0E64F47A31FCCD3BA83C509",1) | Changes has been done. Because we have 15V pin from power card.
| J18 | 2025-04-14 | Align J18 connector properly. |  | Changes has been done.
| L4 | 2025-04-14 | Need to add ferrite bead between DGND and AGND. | =DISPIMG("ID_E909EA5A4FDD4499A186A4081481F23D",1) | Changes has been done.
| R85,R67 | 2025-04-14 | Keep R85 and R67 as zero ohm | =DISPIMG("ID_6C71ADB2F4B54669AAB6CDA90E22D5E8",1) | Have to set manually.
| R21,R23 | 2025-04-14 | Set R21 and R23 to 20K. | =DISPIMG("ID_015AA6EA26C844EA921F8A9D1FD76811",1) | Changes has been done.
| J15 | 2025-04-14 | Verify J15 footprint. |  | Changes has been done. L4 changed to R112 which is 0 ohm.
| L4 | 2025-04-14 | Keep L4 as zero ohm | =DISPIMG("ID_C9136C83582F4A11B97FF33822DD4A7B",1) | Changes has been done.
| J14 | 2025-04-14 | J14 connector +5V supply trace width need to increase | =DISPIMG("ID_432AC574255D480B9AFDCD0517F381C3",1) | Changes has been done. Updated on controller page.
| U6 | 2025-04-14 | Change U6 package to small SOIC, MFR PN: MAX3232CDR. | =DISPIMG("ID_D4699378D1894BA9A734C03AB5522D91",1) | Changes has been done.
| R108 | 2025-04-14 | Keep the value of R108 as 100K ohm. | =DISPIMG("ID_1A47546BB4584A4FB482A4D2E593072E",1) | Changes has been done.
| J10 | 2025-04-14 | Add power supply marking for J10 connector on the silkscreen. | =DISPIMG("ID_D37F93B8A6BF4D94995AF4186CADFBFB",1) | Changes has been done in silkscreen.
| J19 | 2025-08-21 | Changed the power supply connectors (replacing 2 connectors from earlier versions into single connector) | =DISPIMG("ID_C40B8C82DDF44D3B9927FAD266FAF7CB",1) | Changes has been done.
| J11 | 2025-08-21 | Added a separate connector near J6 (Pipette connector) for extra peripherals to add in future. | =DISPIMG("ID_8B80AEFA9BF14FCD9EB3C50696B6BC72",1) | Changes has been done.
| J9 | 2025-08-21 | Changed connector for new RFID module. | =DISPIMG("ID_F1A25E0C8C6E4EE7A621054FAADB48B8",1) | Changes has been done.
SOM Change
| SB127,SB125,SB133,SB132 |  | Remove SB127,SB125,SB133,SB132 | =DISPIMG("ID_15B0F0D79556480C80A48C8E5C526F3C",1)
| SB13,SB160,SB164,SB178,SB181,SB182,SB183 |  | Remove SB13,SB160,SB164,SB178,SB181,SB182,SB183 | =DISPIMG("ID_36814CA93F8A4A6EB7B8966FDEEAB2C2",1)
| JP6, JP7 |  | Remove JP6, JP7 | =DISPIMG("ID_CCDD50D8DFA7414B8FC61639EEC8334D",1)
| JP3 Vin_5 |  | JP3 Vin_5 jumper always keep connected
| X3 |  | X3 need to solder MFR PN- NX3225GD-8.000M-EXS00A-CG04874 | =DISPIMG("ID_F80A23843E8A457D88D02E5FDAAE5112",1)
| C37,C38 |  | Solder C37,C38 10pf/4.3 PF
| SB8,SB9 |  | Solder SB8,SB9 zero Ohm res
Disconnect the Nucleo programmer.
<<INDEX

## Visual Inspection
<<INDEX
Visual Inspection
| Test Group | CQ2BB1
| Test Ref No. | CQ2BB1
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | NA
| Software version | NA
| Test Name | Visual Inspection
| Test Case Objective | Visual verification of PCB to check for Silkscreen.
| Test Case Procedure | Verification of Silkscreen |  |  |  |  | Remark
|  | - All reference designators should be clearly visible. |  |  |  | PASS
|  | - Marking for pin 1 near all IC footprints, connectors and other parts should be visible. |  |  |  | PASS
|  | - All component outlines should be properly aligned and visible after the component is mounted |  |  |  | PASS
|  | - All silkscreen labels should be visible and verify with the silkscreen verification sheet. |  |  |  | PASS
|  | - Polarity marking should be available for the polarized components (Capacitor ,Diode, LED, Connectors) |  |  |  | PASS
|  | - Using Fluke multimeter, measure and note down the resistance between the Supply rails and Ground rails as specified in the Bare Board Impedance Measurement Matrix. |  |  |  | PASS
Bare Board Impedance Measurement Matrix (Expected Result)*
| Test Signals | -12V | +12V | +5V | +3.3V | DGND | Test Result | Remark
| -12V | NA |  |  |  | NA | NA
| +12V | OPEN | NA |  |  | NA | NA
| +5V | OPEN | OPEN | NA |  | NA | NA
| +3.3V | OPEN | OPEN | OPEN | NA | NA | NA
| DGND | OPEN | OPEN | OPEN | OPEN | NA | NA
Bare Board Impedance Measurement Matrix (Expected Result)*
| Test Signals | -12V | +12V | +5V | +3.3V | DGND | Test Result | Remark
| -12V | NA |  |  |  | NA | NA
| +12V | OPEN | NA |  |  | NA | NA
| +5V | OPEN | OPEN | NA |  | NA | NA
| +3.3V | OPEN | OPEN | OPEN | NA | NA | NA
| DGND | OPEN | OPEN | OPEN | OPEN | NA | NA
| Test Case Procedure | Verification of Assembled Components |  |  |  |  | Remark
|  | - All components should be assembled as per the assembly guidelines/BOM. |  |  |  | PASS
|  | - All polarized components should be assembled correctly with proper polarity. |  |  |  | PASS
|  | - All ICs should be oriented as per pin 1 marking in silkscreen. |  |  |  | PASS
|  | - All components should be properly placed on their footprints. |  |  |  | PASS
|  | - There should not be any soldering shorts or solder bridges on PCB |  |  |  | PASS
|  | - There should not be any dry solder points |  |  |  | PASS
| Test Result | PASS
<<INDEX

## Impedance Measurement
<<INDEX
Assembled Board Impedance Measurement
| Test Group | CQ2BB2
| Test Ref No. | CQ2BB2
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | NA
| Software version | NA
| Test Name | Pre power impedance measurements
| Test Case Objective | Verify short / impedance between Supply rails and Ground rails
| Test Case Description | 1. Using Fluke multimeter, measure and note down the resistance between the Supply rails and Ground rails as specified in the Assembled Board Impedance Measurement Matrix. 2. Update the test result
Assembled Board Impedance Measurement Matrix (Expected Result)*
| Test Signals | -12V | +12V | +5V | +3.3V | DGND | Test Result | Remark
| -12V | - | - | - | - | - | NA
| +12V | - | - | - | -
| +5V | - | - | - | - | - | NA
| +3.3V | - | - | - | - | - | NA
| DGND | 0M | 38.7M | 11.2M | 3.74K | - | NA
Assembled Board Impedance Measurement Matrix
| Test Signals | -12V | +12V | +5V | +3.3V | DGND | Test Result | Remark
| -12V | - | - | - | - | - | UNATTEMPTED
| +12V | - | - | - | - | - | UNATTEMPTED
| +5V | - | - | - | - | - | UNATTEMPTED
| +3.3V | - | - | - | - | - | UNATTEMPTED
| DGND | 0M | 38.7M | 11.2M | 3.74K | - | PASS
| Test Result | PASS
<<INDEX
PASS
FAIL
UNATTEMPTED

## Supply Rails
<<INDEX
Supply Rails Measurement
| Test Group | CQ2BB3
| Test Ref No. | CQ2BB3
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | NA
| Software version | NA
| Test Name | Supply Rail and Ripple Measurements
| Test Case objective | To verify tolerance and ripple of supply rail to be within +- 5 %
| Test Case Procedure | Process - - Connect the Board to Back Panel and Power Card. - Turn on +12V to Power Card. - Measure and capture voltage +1.1V w.r.t. DGND - Measure and capture Ripple for + 1.8V w.r.t. DGND - Measure and capture voltage +1.2V w.r.t. DGND - Measure and capture Ripple for +3.3V w.r.t. DGND - Measure and capture voltage +5V w.r.t. DGND - Measure and capture Ripple for +0.85 w.r.t. DGND
| +12V Rail Voltage |  |  |  | +12V Ripple Voltage
| Low Voltage Differential Probe Settings | Coupling | DC |  | Low Voltage Differential Probe Settings | Coupling | AC
|  | Gain | 1 |  |  | Gain | 1
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
| Scope Settings | Coupling | DC |  | Scope Settings | Coupling | AC
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
|  | Trigger | Rising Edge |  |  | Trigger | Rising Edge
| =DISPIMG("ID_6D4CE2CFC5F64A73AA8FCEC16C059EE1",1) |  |  |  | =DISPIMG("ID_8EA428021F70405EA6AFA0A79F405620",1)
| Expected Voltage |  | Observed Voltage |  | Expected Voltage |  | Observed Ripple Voltage
| Test Result | PASS |  |  | Test Result | PASS
| +5V Rail Voltage |  |  |  | +5V Ripple Voltage
| Low Voltage Differential Probe Settings | Coupling | DC |  | Low Voltage Differential Probe Settings | Coupling | AC
|  | Gain | 1 |  |  | Gain | 1
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
| Scope Settings | Coupling | DC |  | Scope Settings | Coupling | AC
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
|  | Trigger | Rising Edge |  |  | Trigger | Rising Edge
| =DISPIMG("ID_53C63F431F8E450DB66412877B6165A9",1) |  |  |  | =DISPIMG("ID_93C28E6BF5D84AEFBFC213CC1F56CA51",1)
| Expected Voltage |  | Observed Voltage |  | Expected Voltage |  | Observed Ripple Voltage
| Test Result | PASS |  |  | Test Result | PASS
| +3.3V Rail Voltage |  |  |  | +3.3V Ripple Voltage
| Low Voltage Differential Probe Settings | Coupling | DC |  | Low Voltage Differential Probe Settings | Coupling | AC
|  | Gain | 1 |  |  | Gain | 1
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
| Scope Settings | Coupling | DC |  | Scope Settings | Coupling | AC
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
|  | Trigger | Rising Edge |  |  | Trigger | Rising Edge
| =DISPIMG("ID_3B1F1F37F3C2468EB1E908D47387CA9A",1) |  |  |  | =DISPIMG("ID_340BA1B69D1F4387AE643EDF2343B5E0",1)
| Expected Voltage |  | Observed Voltage |  | Expected Voltage |  | Observed Ripple Voltage
| Test Result | PASS |  |  | Test Result | PASS
| -12V Rail Voltage |  |  |  | -12V Ripple Voltage
| Low Voltage Differential Probe Settings | Coupling | DC |  | Low Voltage Differential Probe Settings | Coupling | AC
|  | Gain | 1 |  |  | Gain | 1
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
| Scope Settings | Coupling | DC |  | Scope Settings | Coupling | AC
|  | Bandwidth | FULL |  |  | Bandwidth | FULL
|  | Trigger | Rising Edge |  |  | Trigger | Rising Edge
| =DISPIMG("ID_DF00B11E9C28420EACE0F5E74CDAA902",1) |  |  |  | =DISPIMG("ID_05FEC8309FCE4E42AB8E2319FBEE4443",1)
| Expected Voltage |  | Observed Voltage |  | Expected Voltage |  | Observed Ripple Voltage
| Test Result | PASS |  |  | Test Result | PASS
|  |  | Test Result | PASS
<<INDEX

## SWD LED CODE FLASH
<<INDEX
SWD LED CODE FLASH
| Test Group | CQ2BB4
| Test Ref No. | CQ2BB4
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Serial Wire Debugger LED code flash test
| Test Case Objective | To verify functionality of Serial Wire Debugger
| Test Case Procedure | Process - - Use connector J15 for SWD Debugger. - Power On the Board. - Connect SWD debugger (SWDIO, SWCLK, GND, NRST). - Verify MCU detection in the debugger tool or STM32CubeIDE. - Write simple LED blink code. - Reset the MCU after flashing. - Observe LED blinking.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_86505F2351A246D3BAF4F5D8FD255048",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## Temprature Sensor
<<INDEX
Temperature Sensor
| Test Group | CQ2BB5
| Test Ref No. | CQ2BB5
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Temperature Sensor Test
| Test Case Objective | To verify functionality of Temperature Sensor
| Test Case Procedure | Process - - Use connector J13 for Temperature Sensor (CT from J13). - Power On the System Manager Board. - Configure the ADC for 12-bit resolution in the firmware. -Flash the firmware to the Nucleo board. -Reset the board after flashing. -Read ADC values and convert them to temperature. -Verify sensor output with expected temperature range.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_A73FADEF73994847BA0988067C9FF090",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## Display Test
<<INDEX
Display Test
| Test Group | CQ2BB6
| Test Ref No. | CQ2BB6
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Display Test
| Test Case Objective | To verify functionality of Display
| Test Case Procedure | Process - - Use connector J14 for Display. - Power On the System Manager Board. - Connect the 7-inch UART display to the MCU (TX, RX, GND, VCC). -Initialize UART communication in the firmware. -Verify the received data from display. -Receive the data from display and perform LED operation. -Flash the firmware to the Nucleo board. -Reset the board after flashing. -Test touch inputs and verify correct data reception.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_EA26FC03F4E64C6F85204D371C27E2A8",1) | PASS
=DISPIMG("ID_623C35EBAB42466297410B738AE4BB87",1)
|  |  | Test Result | PASS
<<INDEX

## Pipette
<<INDEX
Pipette Test
| Test Group | CQ2BB7
| Test Ref No. | CQ2BB7
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Pipette Test
| Test Case Objective | To verify functionality of Pipette
| Test Case Procedure | Process - - Use connector J6 for Pipette. - Power On the System Manager Board. - Connect the Pipette to Board - Press pipette Buton - Monitor the PC10 pin, if PC10 Pin High Send the puls on nucleo CS4 pin.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | Pipette is working properly. | PASS
|  |  | Test Result | PASS
<<INDEX

## Wifi module
<<INDEX
Wifi Module Test
| Test Group | CQ2BB8
| Test Ref No. | CQ2BB8
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Wifi Module Test
| Test Case Objective | To verify functionality of Wifi Module
| Test Case Procedure | Process - - Power On the System Manager Board. - Initialize communication in the firmware and driver files. - Configure the Wi-Fi module to connect to a network (SSID, password). - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Verify connection by pinging a server or checking an IP address.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_4B1A2B51068045A99380799A600DE1D0",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## BLE
<<INDEX
BLE Test
| Test Group | CQ2BB9
| Test Ref No. | CQ2BB9
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | BLE Test
| Test Case Objective | To verify functionality of BLE
| Test Case Procedure | Process - - Power On the System Manager Board. - Initialize communication in the firmware. - Configure the BLE module in driver files. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Scan and connect using a BLE-enabled device (mobile app or another module). - Verify the connection of BLE and BLE address.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_8F1594E4AAD24941899BB833EC9D8FB8",1) | PASS
=DISPIMG("ID_A5086C6C0E8B4D93850002E4F166A2C1",1)
|  |  | Test Result | PASS
<<INDEX

## USB OTG1 Test
<<INDEX
USB Test
| Test Group | CQ2BB10
| Test Ref No. | CQ2BB10
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | USB OTG1 Test
| Test Case Objective | To verify functionality of USB
| Test Case Procedure | Process - - Use connector J1 for USB OTG1. - Power On the System Manager Board. - Connect the USB OTG adapter and pen drive to the board. - Initialize USB Host mode in the firmware. - Detect the USB device connection. - Mount the file system on the pen drive. - Read and write a test file to verify functionality. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Verify data integrity by checking the file contents.
|  | Expected Result | Observed Result | Test Result
|  |  | =DISPIMG("ID_61661EA0A74946AD8B2473FC4A162C06",1) | PASS
|  | Test Result | PASS
<<INDEX

## USB OTG2 Test
<<INDEX
USB Test
| Test Group | CQ2BB11
| Test Ref No. | CQ2BB11
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | USB OTG2 Test
| Test Case Objective | To verify functionality of USB
| Test Case Procedure | Process - - Use connector J2 for USB OTG2. - Power On the System Manager Board. - Connect the USB OTG adapter and pen drive to the board. - Initialize USB Host mode in the firmware. - Detect the USB device connection. - Mount the file system on the pen drive. - Read and write a test file to verify functionality. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Verify data integrity by checking the file contents.
|  | Expected Result | Observed Result | Test Result
|  |  | =DISPIMG("ID_5B17E7B64BDC4840B1A17D36AF91CDCA",1) | PASS
|  | Test Result | PASS
<<INDEX

## RFID
<<INDEX
RFID Test
| Test Group | CQ2BB12
| Test Ref No. | CQ2BB12
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | RFID Test
| Test Case Objective | To verify functionality of RFID
| Test Case Procedure | Process - - Use connector J9 for RFID. - Power On the System Manager Board. - Connect the RFID module to the MCU - Initialize communication in the firmware. - Send a test command to check module response. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Scan an RFID tag and read its unique ID. - Verify the received data with expected values.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_5888C88F69C8473CBD2D75CEB064CFD2",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## Printer
<<INDEX
Printer Test
| Test Group | CQ2BB13
| Test Ref No. | CQ2BB13
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | v1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Printer Test
| Test Case Objective | To verify functionality of Printer
| Test Case Procedure | Process - - Use connector J7 for Printer Power and J12 for Pinter Connection. - Power On the System Manager Board. -Initialize communication in the firmware. -Send a test print command (e.g., "Hello, World!"). -Flash the firmware to the Nucleo board. - Reset the board after flashing. - Verify that the printer prints the test message correctly.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_DE2ABA55A5404AD6BC72E7FE205E5E6E",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## Ethernet
<<INDEX
Ethernet Test
| Test Group | CQ2BB14
| Test Ref No. | CQ2BB14
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Ethernet Test
| Test Case Objective | To verify functionality of Ethernet
| Test Case Procedure | Process - - Use connector J18 for Ethernet Testing. - Power On the System Manager Board. - Connect the board to a network via an Ethernet cable. - Initialize the Ethernet interface in the firmware. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Obtain an IP address via DHCP or set a static IP. - Ping the board from a PC to verify connectivity. - Test data transmission by sending and receiving packets.
|  | Expected Result | Observed Result | Test Result
|  | Ping from both sides should result in verifying correct link establishment | =DISPIMG("ID_B055495C3FDC43AD916990FEA161AC12",1) | PASS
=DISPIMG("ID_6E23CEE7991247E88E3ACEC285030820",1)
=DISPIMG("ID_3158C83935BC49309808493199235372",1)
|  | Test Result | PASS
<<INDEX

## Heater
<<INDEX
Heater Test
| Test Group | CQ2BB15
| Test Ref No. | CQ2BB15
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Heater Test
| Test Case Objective | To verify functionality of Heater
| Test Case Procedure | Process - - Use connector J5 for Heater. - Power On the System Manager Board. - Initialize GPIO output for heater control in the firmware. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Send a command or set a condition to turn the heater ON. - Verify heater activation (check voltage or temperature rise). - Send a command or set a condition to turn the heater OFF.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_CE9C7081A13D489C8D551AC7DBD39CD9",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## SDRAM
<<INDEX
SDRAM Memory Test
| Test Group | CQ2BB16
| Test Ref No. | CQ2BB16
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | SDRAM Memory Test
| Test Case Objective | To verify functionality of SDRAM Memory
| Test Case Procedure | Process - - Power On the System Manager Board. - Initialize SDRAM controller in the firmware. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Perform a write test (store known data in SDRAM). - Perform a read test (verify stored data matches expected values). - Verify data integrity and SDRAM stability.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_5082A422A79F4C50B42816466746A477",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## QSPI Flash
<<INDEX
QSPI Flash Test
| Test Group | CQ2BB17
| Test Ref No. | CQ2BB17
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | QSPI Flash Test
| Test Case Objective | To verify functionality of QSPI Flash
| Test Case Procedure | Process - - Power On the System Manager Board. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Perform a write test (store known data in QSPI memory). - Perform a read test (verify stored data matches expected values). - Run an erase test and verify memory clearing. - Check data integrity and QSPI performance
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_BAD2BF8D9BDF4A168BBCDAB9CF077FD7",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## ADC
<<INDEX
ADC Test
| Test Group | CQ2BB18
| Test Ref No. | CQ2BB18
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | ADC Test
| Test Case Objective | To verify functionality of ADC
| Test Case Procedure | Process - - Power On the System Manager Board. - Configure SPI settings in the firmware (clock speed, mode, data format). - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Read ADC values from the connected channels. - Convert raw ADC data to voltage values. - Verify readings using a known voltage source or sensor input.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_6CE15B985E1B40309D2EBD2FAAC7CF69",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## IR LED test
<<INDEX
IR LED Test
| Test Group | CQ2BB19
| Test Ref No. | CQ2BB19
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | NA
| Software version | NA
| Test Name | IR LED Test
| Test Case Objective | To verify IR LED working properly
| Test Case Procedure | Process - - Use connector J8 for IR LED Testing. - Power On the System Manager Board. - Expose the photo transistor to different light levels. - Read ADC values and convert them to voltage. - Verify response by observing voltage changes with varying light intensity.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | IR LED is working properly. | PASS
|  |  | Test Result | PASS
<<INDEX

## RS232
<<INDEX
RS232 Test
| Test Group | CQ2BB20
| Test Ref No. | CQ2BB20
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | RS232 Test
| Test Case Objective | To verify functionality of RS232
| Test Case Procedure | Process - - Use connector J3 for RS232 Testing. - Power On the System Manager Board. - Configure UART settings in the firmware (baud rate, parity, stop bits). - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Send a test message from the MCU over RS232. - Receive the message on a PC or another device using a serial terminal. - Verify data integrity by checking if the sent and received messages match.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_F23CB39EAE30448CBDFEEA7409904D80",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## BUZZER
<<INDEX
BUZZER Test
| Test Group | CQ2BB21
| Test Ref No. | CQ2BB21
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Buzzer Test
| Test Case Objective | To verify functionality of Buzzer
| Test Case Procedure | Process - - Power On the System Manager Board. - configure the GPIO as an output in the firmware. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Set the GPIO HIGH to turn the buzzer ON. - Set the GPIO LOW to turn the buzzer OFF. - Verify the buzzer sound activation and deactivation.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | Buzzer is working properly. | PASS
|  |  | Test Result | PASS
<<INDEX

## FAN
<<INDEX
FAN Test
| Test Group | CQ2BB22
| Test Ref No. | CQ2BB22
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | NA
| Software version | NA
| Test Name | Fan Test
| Test Case Objective | To verify Fan working properly
| Test Case Procedure | Process - - Use connector J14 for Fan Testing. - Power ON the System Manager Board. - Connect the FAN.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | Fan is working properly. | PASS
|  |  | Test Result | PASS
<<INDEX

## RTC
<<INDEX
RTC Test
| Test Group | CQ2BB23
| Test Ref No. | CQ2BB23
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.18.0
| Test Name | RTC Test
| Test Case Objective | To verify the functionality of RTC.
| Test Case Procedure | Process - - Power On the System Manager Board. - Connect a backup battery (if required) to the RTC module. - Initialize the RTC peripheral in the firmware. - Flash the firmware to the Nucleo board. - Reset the board after flashing. - Set the current date and time in the RTC. - Read and display the RTC time. - Power cycle the board and verify if the RTC maintains the correct time.
|  |  | Expected Result | Observed Result | Test Result
|  |  |  | =DISPIMG("ID_10D7CEE562B5450DA7A232D26C7FA72F",1) | PASS
|  |  | Test Result | PASS
<<INDEX

## Encryption IC
<<INDEX
Encryption IC
| Test Group | CQ2BB24
| Test Ref No. | CQ2BB24
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Encryption IC Test
| Test Case Objective | To verify the functionality of Encryption IC
| Test Case Procedure | Process - - Power On the Board. - Initialize I²C peripheral on MCU (pins connected to ATSHA204A SDA & SCL). - Send I²C wake-up sequence to ATSHA204A. - Verify device acknowledgment on I²C bus. - Send “Read Serial Number” command to ATSHA204A. - Confirm valid serial number response is received. - Execute “Random Number” command to verify proper operation. - Check response data format and length (matches datasheet spec). - Reset the MCU after test execution. - Confirm ATSHA204A responds correctly after reset.
|  |  | Expected Result | Observed Result | Test Result
UNATTEMPTED
|  |  | Test Result | UNATTEMPTED
<<INDEX

## J11
<<INDEX
Connector J11 Test
| Test Group | CQ2BB25
| Test Ref No. | CQ2BB25
| Test Engineer | Harshit Shakya
| Test Date | 2025-08-25
| Hardware Details | MD_MB_CLQTnew_V3
| Firmware Version | V1.0
| Software version | STM32CubeIDE Version: 1.19.0
| Test Name | Connector J11 Test
| Test Case Objective | To verify the connector J11 is connected with MCU properly
| Test Case Procedure | Process - - Use connector J11 (3-pin GPIO connector). - Power On the Board. - Configure 3 GPIO pins in firmware as digital outputs. - Connect DMM to the 3 GPIO pins and GND. - Toggle each GPIO pin HIGH → LOW in sequence with delay. - Verify voltage transitions on all 3 pins using measurement tool. - Reset the MCU after test execution. - Confirm all GPIO pins return to default LOW state.
|  |  | Expected Result | Observed Result | Test Result
UNATTEMPTED
|  |  | Test Result | UNATTEMPTED
<<INDEX